SE Speeds (v 2.2) – a Users’ Guide

SE Speeds is an easy-to-use utility for setting up a Sedimentation Equilibrium (SE) experiment in an Analytical Ultracentrifuge (AUC).  You enter basic data values – known or estimated – for your solute (molecular weight, partial specific volume) plus basic physical parameters (rotor speed, radial positions of meniscus and cell base, temperature, solvent density) and receive both visual feedback of the predicted equilibrium distribution and a printout of (i) the minimum time which will be taken to attain equilibrium; (ii) the concentration values predicted at the meniscus and at the cell base; (iii) the ratio of the latter two quantities; (iv) the ‘hinge radius’ i.e. point of conservation of cell loading; (v) the sample volume loaded

SE Speeds has two pages, one for ABSORPTION optics, another for INTERFERENCE optics.  The former makes it easy to see if the Beer’s Law is being exceeded at any point in the scan record: the latter gives a warning if the steepness of the fringe gradient at 95% of column length exceeds a value where Svensson anomalies may cause the fringe increment to depart from simple linearity in solute concentration

Use of SE Speeds is not intended as a general guide to the use of the AUC-SE methodology. There are many published reviews under this heading. Equally, the display offered is simple, and for more sophisticated modelling and simulation procedures then use of software such as Peter Schuck’s SEDPHAT is appropriate

How to enter values into SE Speeds

The cells into which you need to enter data are shown in purple. Find the radial positions from the raster display – or assume values if the experiment has not commenced. Results are highlighted in green. The graphical display automatically updates. That’s it!

Novel features in SE Speeds

The ‘time to equilibrium’ feature is based on equations used in older software (Biomols 2). Technically, the translational diffusion coefficient should be used – but this is seldom known. The use of the molecular weight instead (assuming ‘average globular’ physical properties) works fine – for globular proteins. SE Speeds allows for an estimated frictional ratio (as given by SEDFIT on SV data) to be input, and this ratio is used to provide a ‘correction’ to the ‘time to equilibrium’  value computed from molecular weight. The warning for excess fringe gradient is based upon the Svensson equation and work by Peter Lloyd.

Compatibility

This software is written in ‘compatibility mode’ so should run on any version of MS-Excel 97/98 onwards, Windows or Mac (OS-X).

Feedback? – please, to arthur.rowe@nottingham.ac.uk

A few FAQs

- what sort of plot are we looking for? The point has often been made that the information lies ‘in the curvature of the plot’, since every solute component contributes an exponential function to the total plot. However, care must be taken to avoid artefacts arising near the base of the cell, either from Beer’s Law limits being exceeded (absorption optics) or from Svensson anomalies (interference optics)

- is pelleting of solute components a potential problem? If you are trying to analyse the ‘whole solute system’ e.g. to determine average weight values, then this is not really possible if the range in sigma values is great than about 10-fold. Keeping the column fairly short (<2mm) can be helpful

- how short a column can I work with?  With interference optics, going down to 1mm or even less can yield enough data for ‘quality analysis’. But make sure no bits of gasket are waving around near the base!

- is solute concentration an important factor to consider? Yes, it is. SE Speeds is aimed chiefly at experiments carried out in ‘dilute’ solution, roughly < 3 mg/ml for ‘globular’ proteins, < 1 mg/ml for more extended species. It will serve simply as a ‘rough guide’ for experiments in higher concentrations, or for those involving establishment of chemical equilibria

- can I use SE Speeds just to find the sigma value for a given set of MW plus related parameters?  Yes, of course